## Methane analysis
Summary of methods used for methane analysis by laboratories participating in the intercomparison
|  |  | Sample analysis | [10 empty] | Gas standards used by lab groups and comparison with SCOR standards
Lead Scientist | Institute | Reference | Brief description of method | Method for controlling temperature | Method for removal of water vapor | Typical equilibration temperature (oC) | Gas used to equilibrate or purge seawater sample | Water volume ml) | Headspace Volume (ml) | Oven temperature (oC) | Carrier Gas | Analyzer | Column | Personal Standard | Nominal SCOR ARS concentration stated in WG#143 Technical report | Nominal SCOR WRS concentration stated in WG#143 Technical report | Calculated concentration of ARS SCOR standard based on lab standard | Calculated concentration of WRS SCOR standard based on lab standard | Ratio of ARS Calculated concentration to Nominal concentration | Ratio of WRS Calculated concentration to Nominal concentration
Gui-Ling Zhang | Ocean University of China, P.R. China | Zhang et al (2004) J. Geophy. Res. 109 (C7), C07011 | Purge and Trap (sparged with UHP N2,dried, and trapped on Porapak Q trap immersed in LN2) | Laboratory room temperature | Nafion drier and Drierite | 20-25°C | Nitrogen @ 80 ml/min for 5 min | n/a | n/a |  |  | Shimadzu GC14B | 3 m x 3 mm s/s column packed with 80/100 mesh Porapak Q | National Institute of Metrology, China 4.14 ppm +/-2% in nitrogen | SCOR ARS426320 1.938 ppmv | SCOR WRS464573 4.556 ppmv | SCOR ARS426503 1.863 ppmv | SCOR ARS464573 4.539 ppmv | 0.94737 | 0.99627
Sam Wilson | University of Hawaii, USA | Wilson et al (2017) Geophy. Res. Letts. 44: doi: 10.1002/2017GL074458. | Purge and Trap (sparged with UHP He, dried, and trapped on Porapak Q trap immersed in LN2) | Laboratory room temperature | Nafion drier and Drierite | 21 | Helium @ 100 ml/min for 10mins | n/a | n/a | 30 | Helium | Agilent GC7890 | 30 m x 0.32 mm GS-CarbonPLOT capillary column | Scott-Marrin JA01155 20.15 ppm +/-2% in nitrogen | SCOR ARS416388 1.924 ppmv | SCOR WRS460867 4.540 ppmv | Not applicable | Not applicable | Not applicable | Not applicable
Philippe Tortell | University of British Columbia, Canada | Capelle et al (2015) LO Methods 13(7), 345-355 | Purge and Trap (sparged with UHP He, dried, and trapped on Porapak Q trap immersed in LN2) | Laboratory room temperature | Nafion drier | 21 | Helium @ 50 ml/min for 4 mins | n/a | n/a |  |  | Shimadzu QP2010 quadrupole mass spectrometer | 30 m CarbonPLOT column | 10.3 ppm +/-5% in nitrogen | Not applicable | SCOR WRS464565 4.736 ppmv | Not applicable | SCOR WRS464565 4.744 ppmv | Not applicable | 1.0017
Mercedes de la Paz | Instituto de Investigaciones Marinas-CSIC, Spain | de la Paz, (2015) Prog. Oceanogr., 138 (A), 18-31 | Static headspace (19.7 ml N2 in 120 mL bottle, shaken and equilibrated overnight at thermostatted lab temperature. 18 ml of headspace, subsampled by replacement with brine, and 1 ml sampling loop injected onto column | Lab room temperature stable at 20 ºC | Magnesium perchlorate | 20 | Nitrogen | 220 | 19.7 | 60 | Nitrogen | Agilent GC7890 | 2 m 80/100 mesh Porapak Q | NOAA GMD 1.897 ppm, 3.00 ppm, 4.87 ppm | SCOR ARS426347 1.959 ppmv | SCOR WRS464537 4.827 ppmv | SCOR ARS426347 1.939 ppmv | SCOR WRS464565 4.862 ppmv | 0.99 | 0.99
Hermann Bange | GEOMAR, Germany | Arévalo-Martínez et al (2017) JGR Oceans 122 171-184 | Headspace (20 ml bottle with 10 ml of air or helium added, shaken for 20 secs, and then a minimum of 2 h equilibration).Subsample injected using gas-tight syrnge | recording of ambient temperature during headspace removal, thermometer placed in sample vial filled with water | GC injection port equipped with water trap filled with Sicapent | 20-25°C | Helium (99.999%) | In dependence of the expected methane concentrations: 10.3 mL or 50,7 mL | 10 | 60 | Helium | Agilent 5890 | 1.83 m x 3.2 mm packed mole sieve 5A s/s column | Std 4 (Deuste-Steininger, calibrated against NOAA GMD by MPI, Jena) 1.967 ppm +/-1% in synthetic air; Std 5B (Deuste-Steininger): 101.9 ppm +/-2% in synthetic air | SCOR ARS416396 1.940 ppmv | SCOR WRS460872 4.585 ppmv | SCOR ARS416396 1.876 ppmv | SCOR WRS460872 4.612ppmv | 0.967 | 1.006
Cliff Law | NIWA, New Zealand | Weller et al. (2013) Prog. Oceanography 116:193-206 | Headspace equilibration | [8 empty] | Agilent GC6890 |  | CL5 1.751 ppm in nitrogen Std9 5.318 ppm in nitrogen | SCOR ARS416439 1.932 ppmv | SCOR WRS464562 4.552 ppmv | SCOR ARS416439 1.988 ppmv | SCOR WRS464562 4.556 ppmv | 1.029 | 1.0009
Laura Farias | Universidad de Concepción, Chile | Farias et al (2009) Bioegoesciences 6, 3053-3069 | Static headspace (20 ml bottle with 5 ml helium headspace introduced. Shaken and then equilibrated at 30oC for 60 min) 0.5 ml injected into GC |  |  | 30 | Helium | 15 | 5 | 60 |  | Agilent GC6850 | Mole sieve 5A PLOT capiliary column | NOAA GMD 1.875 ppm Air Liquide 1 ppm +/- 5% Air Liquide 5 ppm +/- 5% | SCOR ARS426350 1.953 ppmv | SCOR WRS460877 4.560 ppmv | SCOR ARS426350 2.074 ppmv | SCOR WRS460877 3.963 ppmv | 1.062 | 0.86908
Rob Goddard | Newcastle University, UK | Upstill-Goddard et al (1996), Deep Sea Res, 43, 1669-1682 | Calibrated air headspace created in sample flask and equilibrated by pumped circulation through a closed GC circuit for 20 min. Automated switching of 1 ml sample loop into GC FID carrier gas line, folowe by automatic injection of ambient air and calibration gas samples. | Water bath |  | 25 | Calibrated compressed air (Air Products) | 986 | 96.5 | 60 | UHP nitrogen | Shimadzu GC 14 B -FID | Packed column: Porapak Q 80-100 mesh | SCOR ARS416408 1.961 | SCOR ARS416408 1.961 ppmv | SCOR WRS460877 4.794 ppmv | Not applicable | Not applicable | Not applicable | Not applicable
John Pohlman | USGS, USA | None provided | Headspace: Seawater collected in a 1 L syringe, and 120 ml headspace added. Shaken for 2 mins | Laboratory room temperature | None | 20-25°C | Carbon-free air | 880 | 120 |  |  | Picarro G2201-i cavity ring-down spectrometer |  |  | n/a | n/a
Gregor Rehder | Leibniz-Institute for Baltic Sea Research, Germany | None provided | Purge and Trap; Sparged with UHP N2,dried, and trapped on HayeSep D (60/80 mesh) trap | Laboratory room temperature |  | 20-25°C | Helium | 20 ml |  | 45 | Helium | Shimadzu GC-2014 | HayeSep Q, mesh 80/100, 2 m x 1/8” x 2 mm, stainless steel | 18.47 ppm (± 2 %), 1.88 ppm (± 5 %) (Linde Group, Berlin) | SCOR ARS426360 1.957 ppmv | SCOR WRS464538 4.698 ppmv
Andy Rees | Plymouth Marine Laboratory, UK | Upstill-Goddard et al (1996), Deep Sea Res, 43, 1669-1682 | Calibrated air headspace created in sample flask and equilibrated on orbital shaker for 20 minutes. Headspace removed into gastight syringe and manually injected into GC FID carrier gas line, followed by automatic injection of ambient air and calibration gas samples. | Water bath | Nafion and Magnesium perchlorate column | 25 | Calibrated compressed air (BOC) | 945 | 250 | 55 | BIP N2 (Air Products) | Shimadzu GC-2010 | Hayesep D, 4.0 m length, mesh 60/80 | NOAA - CB11267; 2024.79 ppb Air Products: 0.999, 2.00, 2.99 ± 5% | SCOR ARS416389 1.959 ppmv | SCOR WRS464569 4.762 ppmv | Not determined | Not determined | Not determined | Not determined
Alberto Vieira Borges | University of Liège, Belgium | Borges et al (2015) Nature Geoscience, 8, 637-642 | Headspace (50 ml N2 in 240 ml bottle, overnight equilibration at 21oC 15 ml subsampled, replaced with brine, 5 ml sampling loop injected onto column | [8 empty] | SRI 8610C | Hayesep D, 5.0 m length, mesh 80/100 | 1.045 ± 0.021, 10.23 ± 0.20 in nitrogen | n/a | n/a
Macarena Burgos | Facultad de Ciencias del Mar y Ambientales (CASEM), Spain | None provided | Headspace equilibration in 50 ml SGE syringe. 40 ml withdrawn from sample bottle and 10 ml of nitrogen. Shake for 5 mins using a shaker and wait 5 mins.Inject through a 1 ml loop for methane | Laboratory room temperature | None | 25 | Nitrogen | 40 | 10 | 300 | Nitrogen | Bruker 450-GC | 80/100 Hayesep N and 1.0 m × 1/8-in Molecular Sieve Porapak QS | 1.9514 ppm, 10.06401 ppm | n/a | n/a