## Worksheet 1
Function | Operation | Ancillary field sender ID | Ancillary field command no.
RTC.read | The PLCs are equipped with a real-time clock (RTC) chip and battery to keep track of the date. Once set on commissioning, RTC.read returns the current date and time.
This function loops through each sensor connected on the RS485 bus. Each mesocosm has two sensors (O2 and conductivity–salinity), so each PLC has six sensors connected on its bus.
readMBSensors | O2 sensors have addresses ranging from 10 to 12 for mesocosms0 to 2 of the scenario, respectively.
PC4E sensors have addresses ranging from 30 to 32 for mesocosms0 to 2 of the scenario, respectively.
Sensors are requested individually and in sequence.A request is made every 200 ms.
webSocket.loop | This is a callback function responsible for dealing with the WebSocket communication. The head PLC is the WebSocket server. It listens to branched PLCs' requests and to the monitoring PC requests. Requests are JSON formatted. They always contain ancillary fields: senderID (ID of the entity sending the request), condID (ID of the requested entity), and command (command type of the request). They can also optionally contain a “time” field: Unix-like timestamp (number of seconds since 1 January 1970). | Head PLC (ID = 0)Branched PLCs (ID = 1–3)Monitoring PC (ID = 4) | Request params: set points, PID settings (no. = 0); Request data: measurement values, regulation outputs (no. = 1); Send Params: response to a “Request params” request (no. = 2); Send data: response to a “Request data” request (no. = 3); Calibrate sensor: request to calibrate sensor to a specified value (no. = 4); Request head data: specific data measured by head PLC (pressure and flow rates) (no. = 5); Send head data: a response to a “Request head data” request (no. = 6)
regulationTemperature | This function is responsible for the temperature regulation of the mesocosm. It sets the corresponding three-way valve position using a 0–10 V analog signal. The function first checks if the regulation is in “Manual override” mode. If so, it applies the override set point. If not, it reads the temperature measure in the mesocosm, compares it with the set point, and uses the PID settings to set the valve position.
checkMesocosmes | This functions loops through every mesocosm every 200 ms and reads analog signals (i.e., flow rates and pressure readings).
regulationPressionOnly for HEAD PLC | This function is responsible for the pressure regulation of the mesocosm. It sets the corresponding three-way valve position using a 0–10 V analog signal. The function first checks if the regulation is in “Manual override” mode. If so, it applies the override set point. If not, it reads the pressure measure in the mesocosm, compares it with the set point, and uses the PID settings to set the valve position.
printToSDOnly for HEAD PLC | The head PLC is equipped with a microSD card on which data from all mesocosms are logged every 5 s in one CSV file per day. This is for security only, as the microSD card is not easy to remove from the PLC casing. It should not be removed before the end of the experiment.
regulationSaliniteOnly for branchedPLCs | This function is responsible for the salinity regulation of the mesocosm. It sets the corresponding three-way valve position using a 0–10 V analog signal. The function first checks if the regulation is in “Manual override” mode. If so, it applies the override set point. If not, it reads the salinity measure in the mesocosm, compares it with the set point, and uses the PID settings to set the valve position.