## Nitrous oxide
Summary of methods used for nitrous oxide analysis by laboratories participating in the intercomparison
Gas standards used by lab groups and comparison with SCOR standards
Lead Scientist | Institute | Reference | Brief method of Gas extraction/equilibration | Typical equilibration temperature (oC) | Gas used to equilibrate or purge seawater sample | Water volume ml) | Headspace Volume (ml) | Oven temperature (oC) | Carrier Gas | Column | Analyzer | Personal lab standard | Nominal SCOR ARS concentration stated in WG#143 Technical report | Nominal SCOR WRS concentration stated in WG#143 Technical report | Calculated concentration of ARS SCOR standard based on lab standard | Ratio of ARS Calculated concentration to Nominal concentration
Gui-Ling Zhang | Ocean University of China, P.R. China | Walter et al. (2005) 32, L23613, doi: 10.1029/2005GL024619 | Headspace equilibration | [6 empty] | 3 m x 3 mm s/s column packed with 80/100 mesh Porapak Q | Shimadzu GC14B | National Institute of Metrology, China 230, 610, 3040 ppb +/-2% in nitrogen | SCOR ARS426320 326.0 ppb | SCOR WRS464573 22326 ppb | SCOR ARS426320 326.8 ppb | 1.0025
Sam Wilson | University of Hawaii, USA | Wilson et al (2017) Geophy. Res. Letts. 44: doi: 10.1002/2017GL074458. | Purge and Trap (sparged with UHP He, dried, and trapped on Porapak Q trap immersed in LN2) | 21 | Helium @ 100 ml/min for 10mins | n/a | n/a | 30 | Helium | 30 m x 0.32 mm GS-CarbonPLOT capillary column | Agilent GC7890 | Scott-Marrin JA01155 4810 +/-2% in nitrogen, NOAA GMD CB10298 357.56 ppb in air | SCOR ARS416388 325.3 ppb | SCOR WRS460867 22271 ppb | SCOR ARS416388 336.6 ppb | 1.0347
Philippe Tortell | University of British Columbia, Canada | Capelle et al (2015) LO Methods 13(7), 345-355 | Purge and Trap (sparged with UHP He, dried, and trapped on Porapak Q trap immersed in LN2) | 21 | Helium @ 50 ml/min for 4 mins |  |  |  |  | 30 m CarbonPLOT column | Shimadzu QP2010 quadrupole mass spectrometer | 4000 ppb +/-5% in nitrogen |  | SCOR WRS464565 23401 ppb
Mercedes de la Paz | Instituto de Investigaciones Marinas-CSIC, Spain | de la Paz, (2015) Prog. Oceanogr., 138 (A), 18-31 | Static headspace (19.7 ml N2 in 120 mL bottle,shaken and equilibrated overnight at thermostatted lab temperature. Then, 18 ml of headspace subsampled by replacement with brine, and 1 ml sampling loop injected onto column | 20 | Nitrogen | 220 | 19.7 | 60 | Nitrogen | 4 m 80/100 mesh Porapak Q | Agilent GC7890 | NOAA GMD 324.97 ppb, 1020, 3100 | SCOR ARS426347 332.1 ppb | SCOR WRS464537 23885 ppb | SCOR ARS426347 330.2 ppb | 0.9943
Hermann Bange | GEOMAR, Germany | Arévalo-Martínez et al (2017) JGR Oceans 122 171-184 | Headspace (20 ml bottle with 10 ml of air or helium added, shaken for 20 secs, and then a minimum of 2 h equilibration) | 20-25°C | Helium (99.999%) | 10.3 | 10 | 190 | Argon/Methane (95% /5%) | 1.83 m x 3.2 mm packed mole sieve 5A s/s column | Agilent 5890 | Std4 (Deuste-Steininger, calibrated against NOAA GMD Std by MPI, Jena): 318.2 +/-1% in synthetic air Std 5B (Deuste-Steininger): 994.1 ppb +/- 2% in synthetic air manual dilutions of Std 5B with helium (1:2, 2:1,average precision: 3%) | SCOR ARS416396 328.7 ppb | SCOR WRS460872 22627 ppb | SCOR ARS416396 322.85 ppb | 0.9822
Cliff Law | NIWA, New Zealand | None provided | Headspace equilibration | [7 empty] | Agilent GC6890 | CL5 323.11 ppb in nitrogen Std9 1010.29 ppb in nitrogen | SCOR ARS416439 325.9 ppb | SCOR WRS464562 22367 ppb | SCOR ARS416396 319.6 ppb | 0.9807
Laura Farias | Universidad de Concepción, Chile | Farias et al (2009) Bioegoesciences 6, 3053-3069 | Static headspace (20 ml bottle with 5 ml helium headspace introduced. Shaken and then equilibrated at 30oC for 60 min) 0.5 ml injected into GC | 30 | Helium | 15 | 5 | 60 |  | 3-m Poropak | Varian 3380 | Air Liquide 250ppb +/- 5% Air Liquide 500ppb +/-5% Air Liquide 1000ppb +/-5% | SCOR ARS426350 329.7 ppb | SCOR WRS464553 22746 ppb | SCOR ARS426350 335.5 ppb | 1.0176
Rob Goddard | Newcastle University, UK | Upstill-Goddard et al (1996), Deep Sea Res, 43, 1669-1682 | Calibrated air headspace created in sample flask and equilibrated by pumped circulation through a closed GC circuit for 20 min. Automated switching of 1 ml sample loop into GC FID carrier gas line, folowe by automatic injection of ambient air and calibration gas samples. | 25 | Calibrated compressed air (Air Products) | 986 | 96.5 | 60 | UHP Nitrogen | Q chromatographic column | Shimadzu GC 14 B -ECD | SCOR ARS416408 331.7 | SCOR ARS416408 331.7 ppb | SCOR WRS464579 22746 ppb | Not applicable | Not applicable
Gregor Rehder | Leibniz-Institute for Baltic Sea Research, Germany | None provided | Purge and Trap; Sparged with UHP N2,dried, and trapped on HayeSep D (60/80 mesh) trap | 20-25°C | Helium | 20 ml |  | 45 | Helium | HayeSep Q, mesh 80/100, 2 m x 1/8” x 2 mm, stainless steel | Shimadzu GC-2014 | Matheson 100 ppb Matheson 1000 ppb | SCOR ARS426360 326.5 ppb | SCOR WRS464538 23244 ppb | SCOR ARS426366 293.9 ppb | 0.9002
Alberto Vieira Borges | University of Liège, Belgium | Borges et al (2017) Ecosystems, doi: 10.1007/s10021-017-0171-7 | Headspace (50 ml N2 in 240 ml bottle, overnight equilibration at 21oC after initial shaking | [6 empty] | Hayesep D, 5.0 m length, mesh 80/100 | SRI 8610C | 210 ± 21, 1990 ± 100 in nitrogen
Andy Rees | Plymouth Marine Laboratory, UK | Upstill-Goddard et al (1996), Deep Sea Res, 43, 1669-1682 | Calibrated air headspace created in sample flask and equilibrated on orbital shaker for 20 minutes. Headspace removed into gastight syringe and manually injected into GC FID carrier gas line, followed by automatic injection of ambient air and calibration gas samples. | 25 | Calibrated compressed air (BOC) | 945 | 250 | 55 | BIP N2 (Air Products) | Poropak Q, 2.0m length, mesh 80/100 | Shimadzu GC-2010 | NOAA - CB11267; 359.73 ppb Air Products: 301.0, 400.0, 500.0 ± 1% ppb | SCOR ARS416389 330.7 ppb | SCOR WRS464569 23578 ppb | Not Determined | Not Determined
Macarena | Facultad de Ciencias del Mar y Ambientales (CASEM), Spain | None provided | Headspace equilibration in 50 ml SGE syringe. 40 ml withdrawn from sample bottle and 10 ml of nitrogen. Shake for 5 mins using a shaker and wait 5 mins | 25 | Nitrogen | 40 | 10 | 350 | Ar/CH4 (95%/5%) | 0.5 m × 1/8-in stainless steel Hayesep N column (80/100) and 2.0 m × 1/8-in stainless steel Hayesep D column. | Bruker 450-GC | 304.61 ppb, 474.38 ppb
Alyson Santoro | University California, Santa Barbara, USA | None provided | Headspace (35 mL UHP N2 in 160 mL bottle,10 mL overpressurization with UHP N2, equilibration at RT with gentle shaking for a minimum of 2h. 5 mL headspace withdrawl, injection into 1 mL sample loop. | 22.5 | UHP N2 | 160.7 | 35 | 100 | UHP N2 | Packed column, split; 2 x 3m, Haysep D | SRI 8610C, ECD temp 353ºC | Matheson 0.1 ppm and Matheson 1 ppm | SCOR ARS426366 326.5 ppb | SCOR WRS464585 22549 ppb | 401.5 ppb | 1.213
John Bullister | NOAA PMEL | Bullister and Wisegarver (2008) Deep-Sea Res., 55, 1063-1074, 2008. | Purge and Trap